## Instructions for Use
AAFS Standards Resources and Training
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions, an on-demand training webinar and instructional videos.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASTM E3189-19
ANSI/ASTM E3189-19
Standard Practice for Separation of Ignitable Liquid Residues from Fire Debris Samples by Static Headspace Concentration onto an Adsorbent Tube
Discipline: | Fire Debris & Explosives
Checklist version: | V1
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Apparatus | 6 | Section Title
Static Headspace Concentration Sampling Apparatus | 6.1 | Section Title
Static Headspace Concentration Sampling Apparatus | 6.1.1 | Requirement | Sampling Device A device capable of concentrating volatile compounds in air, consisting of an adsorbent tube with the outlet connected to an air-tight syringe or pump and the inlet connected to a needle or plastic tip.
Static Headspace Concentration Sampling Apparatus | 6.1.2 | Requirement | Puncturing Device A device such as a pin or a nail that is capable of puncturing a small hole in the sample container. The size of the hole is such that a syringe needle or plastic pipette tip can fit (loosely) into the opening.
Static Headspace Concentration Sampling Apparatus | 6.1.3 | Requirement | Sampling System A system that consists of a sampling and puncturing device, capable of withdrawing a portion of headspace through a hole in the container and concentrating volatile compounds from this headspace portion onto an adsorbent tube.
Static Headspace Concentration Sampling Apparatus | 6.1.3.1 | Requirement | The sampling system must be designed and used in a manner that prevents the loss of volatile compounds from the container and contamination of the sample from the apparatus itself or the laboratory environment.
Apparatus | 6.2 | Requirement | Thermal Desorption System A system capable of desorbing the trapped volatile compounds from an adsorbent tube by means of elevated temperature, refocusing them on a cold-trap and subsequently introducing them to a capillary GC column by flash heating. The desorption apparatus is directly coupled to a GC-MS.
Apparatus | 6.3 | Requirement | Pump A vacuum or hand pump capable of drawing air at a flow rate of at least 2 mL/min to 80 mL/min.
Apparatus | 6.4 | Requirement | Oven An oven large enough to accommodate the sample container, and capable of maintaining the required temperature uniformly throughout.
Apparatus | 6.5 | Requirement | Temperature Measuring Device A thermometer or thermocouple capable of measuring oven temperatures in the required range of operation, to within approximately 5°C.
Reagents and Materials | 7 | Section Title
Reagents and Materials | 7.1 | Requirement | Purity of Reagents Reagent grade or better chemicals shall be used in all tests. Unless otherwise indicated, it is intended that all reagents conform to the specifications of the Committee on Analytical Reagents of the American Chemical Society where such specifications are available. Other grades may be used, provided it is first ascertained that the reagent is of sufficient high purity through evaluation of appropriate blank(s) to permit its use without lessening the accuracy of the determination.
Reagents and Materials | 7.2 | Requirement | Filter Permeable material such as silanized glass wool, wire mesh or cotton held in place with a frit.
Reagents and Materials | 7.3 | Requirement | Solid Adsorbent Activated carbon, Tenax TA, or equivalent.
Reagents and Materials | 7.4 | Requirement | Sampling Tubes Glass Pasteur pipettes or equivalent glass tubes for solvent elution, and stainless steel or glass tubes for thermal desorption.
Reagents and Materials | 7.5 | Requirement | Adsorbent Tubes Sampling tubes packed with a solid adsorbent.
Reagents and Materials | 7.5.1 | Informational Text | Pre-packed adsorbent tubes are commercially available. Empty sampling tubes that can be packed by the purchaser are also available.
Reagents and Materials | 7.5.1.1 | Informational Text | Activated carbon tubes and equivalent for solvent elution can be made by inserting a filter into the bottom of a sampling tube (approximately 5-mm diameter), then adding 2.5 cm to 5 cm of activated carbon or equivalent, which is held in place with a second filter.
Reagents and Materials | 7.5.1.2 | Informational Text | Tenax TA tubes and equivalent for thermal desorption can be made by inserting a filter into one end of a sampling tube, adding approximately 80-mg pre-conditioned Tenax TA or equivalent, and then packing tightly with a second filter. NOTE 1—Tenax TA or equivalent is conditioned by heating, in accordance with instrument manufacturer or supplier instructions.
Reagents and Materials | 7.6 | Requirement | Screw cap or crimp-top glass vials with polytetrafluoroethylene (PTFE) lined seals.
Reagents and Materials | 7.7 | Requirement | Disposable Syringes 1 mL to 100 mL.
Reagents and Materials | 7.8 | Requirement | Disposable syringe needles, tips, or equivalent.
Reagents and Materials | 7.9 | Requirement | Tape, rubber sleeve stopper, or equivalent.
Reagents and Materials | 7.1 | Requirement | Elution Solvent Suitable elution solvents include carbon disulfide, n-pentane, dichloromethane and diethyl ether.
Sample Preparation | 8 | Section Title
Sample Preparation | 8.1 | Requirement | Observe the appropriate procedures for handling and documentation of all submitted samples (see Guide E1459 and Practice E1492).
Sample Preparation | 8.2 | Requirement | Prepare the fire debris sample container for static headspace concentration onto an adsorbent tube.
Sample Preparation | 8.2.1 | Informational Text | The sampling system is designed so that the headspace sample can be extracted from the container in which it was received at the laboratory.
Sample Preparation | 8.2.1.1 | Informational Text | Alternatively, the fire debris sample, or a portion of it, can be transferred to a clean sample container such as a jar, can, or bag suitable for static headspace concentration onto an adsorbent tube.
Sample Preparation | 8.2.1.2 | Requirement | Verify the cleanliness of the transfer container prior to sample transfer. Cleanliness is determined by means of analysis of a static headspace concentration sample obtained from the empty transfer container using the same conditions as will be used for the questioned sample.
Sample Preparation | 8.2.1.3 | Requirement | Allow the transfer container with sample to equilibrate for at least one hour before sampling.
Sample Preparation | 8.2.2 | Requirement | Create a hole in the container with the puncturing device such that a syringe needle or plastic tip can be introduced. The hole is created in the lid of a can or jar, or in the headspace area of a bag.
Sample Preparation | 8.2.2.1 | Requirement | Seal the hole with tape, rubber sleeve stopper, or equivalent. NOTE 2—Cans that are designed for fire debris samples, with a hole pre-fitted with a rubber sleeve stopper, are commercially available.
Sample Preparation | 8.3 | Requirement | When necessary, heat the container of fire debris sample at the selected temperature prior to the headspace sampling, in order to release any ignitable liquid residues as a vapor into the headspace.
Sample Preparation | 8.3.1 | Recommendation | Heating is not recommended if there are other evidentiary considerations such as preservation of DNA or latent fingerprints.
Sample Preparation | 8.3.2 | Requirement | Place the container into an oven for a suitable heating time and temperature, based upon parameters determined during the method verification as outlined in Section 11. NOTE 3—Sampling at room temperature is suitable for the extraction of the more volatile compounds, while heating increases the concentration of the less volatile compounds in the headspace.
Adsorption Procedure | 9 | Section Title
Adsorption Procedure | 9.1 | Requirement | Remove the container from the oven, if heated.
Adsorption Procedure | 9.2 | Requirement | Prepare the container and adsorbent tube for sampling.
Adsorption Procedure | 9.2.1 | Requirement | Remove the ends or caps of the adsorbent tube and connect the outlet of the adsorbent tube to a syringe or pump. Connect the inlet of the adsorbent tube to a needle or plastic tip and insert this needle or plastic tip through the bag or through the hole created in the container lid. Disposable plastic tips or needles are used to protect the outside of the sampling tube from possible contamination due to contact with the inside of the container.
Adsorption Procedure | 9.3 | Requirement | Withdraw a sample of the headspace from the container through the adsorbent tube, using a syringe or pump.
Adsorption Procedure | 9.3.1 | Requirement | Sampling from a heated container removed from an oven is performed without delay while it is still at or near the elevated temperature.
Adsorption Procedure | 9.3.2 | Informational Text | A syringe is typically used for sampling a headspace volume between 1 mL and 100 mL.
Adsorption Procedure | 9.3.3 | Informational Text | A pump is typically used for sampling more than 100 mL of headspace (generally up to a maximum 10 % of the container volume). The required sampling time depends on the pump flow rate.
Adsorption Procedure | 9.3.3.1 | Informational Text | Pumps are typically used when the adsorbent tube will be desorbed by solvent elution because solvent elution dilutes the recovered volatile compounds. There is no dilution with thermal desorption.
Adsorption Procedure | 9.3.4 | Informational Text | A suitable sample volume is determined during the method verification as outlined in Section 11. It can vary from 1 mL to 100 mL or more (generally up to maximum 10 % of the container volume) depending on the selected heating temperature, the size of the container, the concentration of ignitable liquid residues present in the sample, the type and amount of sample material in the container, and the selected adsorbent type and desorption procedure (thermal desorption or solvent elution).
Adsorption Procedure | 9.3.4.1 | Informational Text | An indication of the ignitable liquid residue concentration in a questioned sample can be obtained from a preanalysis at room temperature using the method in this standard, or from an alternative method such as static headspace sampling (Practice E1388) at room temperature or elevated temperature and subsequent analysis with gas chromatographyflame ionization detection (GC-FID) or GC-MS.
Adsorption Procedure | 9.4 | Requirement | Remove the sampling device from the container, seal the opening in the container with a septum, tape, or other means that prevent vapors from escaping or entering; remove the syringe and discard the needle or plastic tip. If the tube is to be thermally desorbed, recap it immediately.
Desorption Procedure | 10 | Section Title
Desorption Procedure | 10.1 | Informational Text | Solvent elution is typically used with activated carbon or equivalent adsorbents.
Desorption Procedure | 10.1.1 | Requirement | Follow the safety precautions described in the safety data sheet (SDS) of the elution solvent that is used.
Desorption Procedure | 10.1.2 | Requirement | Pass approximately 1 mL of elution solvent through the adsorbent tube (containing activated carbon or the equivalent) into a vial. If possible, allow the solvent to remain in contact with the adsorbent material for approximately 5 minutes before expelling. An aspirator bulb can be used if necessary to push the solvent through the tube.
Desorption Procedure | 10.1.2.1 | Requirement | Alternatively, completely remove the adsorbent material from the tube and agitate up to 5 minutes in a container with elution solvent. Transfer the solvent extract into a vial and cap the vial in preparation for analysis.
Desorption Procedure | 10.1.3 | Informational Text | If necessary, the solvent extract can be concentrated by evaporation under a stream of dry inert gas, dry nitrogen, or dry air that is filtered to remove hydrocarbons. (Warning— Evaporation can result in loss of some volatile compounds.)
Desorption Procedure | 10.1.4 | Informational Text | Refer to Practice E2451 for short term and long term storage of ignitable liquid extracts.
Desorption Procedure | 10.2 | Informational Text | Thermal desorption is typically used with Tenax TA or equivalent adsorbents.
Desorption Procedure | 10.2.1 | Requirement | Transfer the adsorbent tube to the thermal desorption system for desorption and subsequent GC-MS analysis.
Desorption Procedure | 10.2.1.1 | Informational Text | In the thermal desorption system, the volatile compounds trapped on the adsorbent tube are desorbed, and refocused onto a cold trap prior to introduction into the GC-MS.
Quality Assurance | 11 | Section Title
Quality Assurance | 11.1 | Requirement | Verify the practice of static headspace concentration onto an adsorbent tube prior to use on questioned samples utilizing Test Method E1618.
Quality Assurance | 11.1.1 | Requirement | Use a selection of ignitable liquids that together cover the entire range of compounds to be identified with this practice, a range of concentrations of those ignitable liquids, and ignitable liquids on matrices such as wood and carpet created to simulate common questioned samples.
Quality Assurance | 11.1.2 | Requirement | Determine suitable sample volumes for the range of ignitable liquid concentrations applicable for this practice.
Quality Assurance | 11.1.2.1 | Informational Text | The sample volume required generally decreases with increased heating temperature and with increased concentration of ignitable liquid residues present in the sample. The required sample volume for thermal desorption is generally less compared to solvent elution, because of the solvent dilution which is inherent in the latter desorption procedure.
Quality Assurance | 11.1.3 | Requirement | Determine the heating system temperature and the time period for heating of the sample container.
Quality Assurance | 11.1.3.1 | Informational Text | Heating is not required to recover low boiling compounds up to C10, but can improve recovery.
Quality Assurance | 11.1.3.2 | Informational Text | Heating is required to recover higher boiling compounds, at temperatures typically between 60°C and 90°C for compounds up to C16, and as high as 130°C for compounds above C20.
Quality Assurance | 11.1.3.3 | Informational Text | A temperature below 90°C is most appropriate for samples that contain moisture or that are wet.
Quality Assurance | 11.1.4 | Informational Text | The heating time can vary from minutes up to several hours, depending on the selected temperature, the size of the container, and the type and amount of matrix in the container.
Quality Assurance | 11.1.4.1 | Informational Text | The required heating time generally increases with increased container volume or increased amount of matrix occupying that volume, or both.
Quality Assurance | 11.2 | Requirement | Determine the maximum volume that can be sampled onto the adsorbent without analyte breakthrough, for common ignitable liquid classes described in Test Method E1618. Do not exceed this volume in practice.
Quality Assurance | 11.2.1 | Requirement | Prepare sample containers with a quantity of analyte mixture or ignitable liquid that is approximately ten times the upper limit of detection. Sample different volumes from each container in the same manner as questioned samples (as described in Sections 8 and 9), but place two adsorbent tubes in series.
Quality Assurance | 11.2.2 | Requirement | Analyze the secondary adsorbent tubes for the presence of ignitable liquid compounds, using the same desorption parameters as will be used for questioned samples (as described in Section 10) and using the same analysis parameters used for questioned samples and in accordance with Test Method E1618.
Quality Assurance | 11.2.3 | Informational Text | Analyte breakthrough during adsorption is possible particularly when a pump is used for sampling. Loss of some analytes can result if the total sample volume is too high.
Quality Assurance | 11.3 | Requirement | Store adsorbent material and capped adsorbent tubes in a clean environment at all times to avoid adsorption of interference compounds.
Quality Control | 12 | Section Title
Quality Control | 12.1 | Requirement | Check the purity of each new bottle of the elution solvent before use, by evaporating a portion of the solvent to at least 50 % of its original volume, or, if concentration of a solvent extract is performed, by evaporating at least twice the volume used in the analysis. Evaporation is performed as described in 10.1.3. Analyze the evaporated solvent using the same conditions as used for questioned samples and in accordance with Test Method E1618.
Quality Control | 12.2 | Requirement | Check each new, cleaned, or reconditioned batch of adsorbent tubes before use for contaminants by desorbing, at a minimum, at least one of the tubes using the same desorption procedure as used for questioned samples (as described in Section 10). Analyze these adsorbent tubes, or the samples prepared from these adsorbent tubes, using the same conditions as used for questioned samples and in accordance with Test Method E1618.
Quality Control | 12.2.1 | Requirement | Activated carbon or equivalent is considered to be disposable, whereas Tenax TA or equivalent is reusable. Check reusable adsorbents batch-wise, at a minimum, after being cleaned or reconditioned, prior to being re-used.
Quality Control | 12.2.1.1 | Informational Text | The lifetime of reusable adsorbents is usually specified by the suppliers and thermal desorption instrument vendors, but can be exceeded in practice if their adsorption and desorption efficiencies are regularly checked as described in 12.4.
Quality Control | 12.3 | Requirement | Check the cleanliness of the sampling system and desorption procedure for contaminants prior to processing a batch of questioned samples, which can include preparing an empty container as a negative control. Prepare and sample the empty container in the same manner as questioned samples (as described in Sections 8 and 9). Desorb the adsorbent tube using the same desorption process and conditions as will be used for questioned samples (as described in Section 10), and analyze using the same conditions used for questioned samples and in accordance with Test Method E1618.
Quality Control | 12.3.1 | Informational Text | A negative control is a clean, empty sample container similar to that used for questioned samples.
Quality Control | 12.4 | Requirement | Check the adsorption and desorption efficiency of the system near the beginning and end of every sequence of questioned samples, or at a frequency established during method verification, by preparing a positive control containing a known volume of ignitable liquid, or of a standard test mixture in accordance with Test Method E1618 that spans a range of boiling points. Prepare and sample the container in the same manner as questioned samples (as described in Sections 8 and 9). Desorb the adsorbent tube using the same desorption process and conditions as will be used for questioned samples (as described in Section 10), and analyze using the same conditions used for questioned samples and in accordance with Test Method E1618.
Quality Control | 12.4.1 | Informational Text | The positive control is used to verify that the adsorption and desorption efficiency of the system is within the required specifications from the verification studies of the practice. The adsorption and desorption efficiency is determined by evaluating the recoveries of selected compounds of interest using peak area or peak height.
Quality Control | 12.4.2 | Informational Text | The positive control includes a periodic check of the adsorption and desorption capabilities of the individual or batch of adsorbent tubes in use.
Quality Control | 12.5 | Requirement | Check empty sampling tubes as in accordance with laboratory policy to demonstrate the cleanliness of the desorption process and the GC-MS. Desorb the empty tubes using the same desorption process and conditions as used for questioned samples (as described in Section 10), and analyze using the same conditions used for questioned samples in accordance with Test Method E1618.
Quality Control | 12.6 | Requirement | If an internal standard is routinely added to questioned samples and reference samples, then also include it in the prepared blank.
Conditioning and Regeneration | 12.7 | Section Title
Conditioning and Regeneration | 12.7.1 | Requirement | Condition Tenax TA or equivalent adsorbents at elevated temperatures before initial use according to the laboratory or manufacturer’s instructions.
Conditioning and Regeneration | 12.7.2 | Requirement | Sampling tubes that indicate deterioration in performance, such as low recovery rate, or nonzero blanks, can be regenerated in the same manner. Check regenerated tubes before re-use as described in 12.4.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed |  | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics