## Instructions for Use
AAFS Standards Resources and Training
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions, an on-demand training webinar and instructional videos.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASTM E2927-16e1
ANSI/ASTM E2927-16e1
Standard Test Method for Determination of Trace Elements in Soda-Lime Glass Samples Using Laser Ablation Inductively Coupled Plasma Mass Spectrometry for Forensic Comparisons
Discipline: | Trace Evidence
Checklist version: | V1
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Apparatus | 6 | Section Title
Apparatus | 6.1 | Requirement | LA-ICP-MS A Laser Ablation system coupled to an ICP-MS instrument is employed. Since there are several manufacturers for both laser ablation units and ICP-MS instruments, the instrument maker, model, configuration, and major operational parameters (that is, laser wavelength for the laser and mass selective detector type for the ICP-MS) of both instruments shall be noted within the analysis results. The most common laser wavelengths used for glass analysis are 266 nm, 213 nm, and 193 nm. Either quadrupole or magnetic sector ICP-MS instruments are suitable for this test method.
Apparatus | 6.2 | Requirement | Prior to the analysis on the day it is used, the ICP-MS shall be tuned according to the manufacturer’s recommendations covering the mass range of the elements to be measured. The instrument shall be adjusted for maximum sensitivity, best precision, and to minimize oxides and doubly charged ion interferences. The use of a glass reference material, such as NIST 612, is recommended during the tuning and performance check. Detector cross calibrations (pulse/analog) shall be performed before any measurements when two detector modes are used in the analysis.
Apparatus | 6.3 | Informational Text | In order to prepare for data acquisition, the signals of the following isotopes are monitored in the ICP-MS; lithium (7Li), magnesium (24 or 25Mg), aluminum (27Al), silicon (29Si), potassium (39K), calcium (42Ca), iron (57Fe), titanium (49Ti), manganese (55Mn), rubidium (85Rb), strontium (88Sr), zirconium (90Zr), tin (118Sn), barium (137Ba), lanthanum (139La), cerium (140Ce), neodymium (146Nd), hafnium (180Hf) and lead (208Pb). This procedure may be applicable to other elements and other isotopes (for example, 206, 207Pb); however, those elements listed above are considered to provide the most discrimination power for soda-lime glass comparisons. Alternatively, other isotopes such as 56Fe may be monitored using ICP-MS with advanced technology to remove interferences (for example, sector field ICP-MS or reaction cells).
Apparatus | 6.4 | Informational Text | Either argon or helium may be used as a carrier gas to transport the particles from the ablation cell to the plasma. The use of helium carrier gas has been reported to result in fewer fractionation effects than the use of argon as a carrier.
Hazards | 7 | Section Title
Hazards | 7.1 | Requirement | Commercial laser ablation units are enclosed type I lasers. However, laser systems typically used for analysis of glass generate high energy radiation that may pose serious risks to eye safety if exposed to the eye. Interlocks shall not be bypassed or disconnected.
Hazards | 7.2 | Requirement | The argon plasma shall not be observed directly without protective eyewear. Potentially hazardous UV light may be emitted.
Hazards | 7.3 | Requirement | ICP-MS instruments generate high amounts of radiofrequency energy in their RF power supply and torch boxes that is potentially hazardous if allowed to escape. Safety devices and safety interlocks shall not be bypassed or disconnected.
Calibration and Standardization | 8 | Section Title
Calibration and Standardization | 8.1 | Requirement | A calibration curve using multiple glass standards or using a single glass standard may be used for quantitation for LA-ICP-MS analysis of glass. Any calibration standard shall be matrix-matched to the sample and well-characterized. The calibration standard(s) shall be traceable to an accepted standard. For glass analysis there are several standards that are available such as the NIST Standard Reference Materials (that is, NIST 610, NIST 612, NIST 614) and the float glass standard glasses (FGS1, FGS2) evaluated by the European group NITECRIME and distributed by the Bundeskriminalamt, Germany. A normalization standard, preferably silicon (29Si), shall be used to normalize the signal. The use of a normalization standard is needed to adjust for differences in ablation yield between the ablated materials. Since silicon is present as a major component in all soda-lime glass (~70 to 72 % as SiO2), a low abundance isotope (29Si) is commonly used as the normalization standard for this method. If this method is used for the analysis of other glass types, the concentration of the normalization standard shall be determined prior to quantitative analysis.
Calibration and Standardization | 8.2 | Requirement | In addition to the calibration standard, at least one additional glass standard reference material shall should be measured with each sample set as a quality control check for the accuracy and precision of the method. The quality control specifications shall be set by each laboratory.
Calibration and Standardization | 8.3 | Requirement | As a minimum, calibration standards are required at the beginning and the end of the analytical sequence in order to adjust for instrument drift over time. Acceptance criteria for the calibration shall be defined by each laboratory and shall include the use of calibration verification standards.
Procedure | 9 | Section Title
Procedure | 9.1 | Informational Text | If necessary, samples may be cleaned to remove any surface contamination by washing or pre-ablation, or both, prior to analysis. Cleaning may include washing samples with soap and water, with or without ultrasonication, and rinsing in deionized water, followed by rinsing in acetone, methanol, or ethanol, and drying. Soaking in various concentrations of nitric acid for 30 minutes or longer, rinsing with deionized water and ethanol, and drying prior to analysis removes most surface contamination without affecting the measured concentrations of elements inherent in the glass. However, the use of nitric acid may remove some surface coatings that may be present.
Procedure | 9.2 | Informational Text | Multiple samples and standards may be placed together in the ablation cell as long as their positions are documented.
Procedure | 9.3 | Requirement | The samples or standards, or both, shall be secured in the ablation cell using double-sided tape or other adhesive. Orient the sample to avoid an original surface of the glass. The known and questioned samples shall be treated equally. The following steps shall be followed:
Procedure | 9.3.1 | Requirement | Purge the ablation cell with the carrier gas between samples to avoid any contamination.
Procedure | 9.3.2 | Informational Text | If pre-ablation cleaning is performed it may be done at this point.
Procedure | 9.3.3 | Requirement | Focus the laser beam at the surface of the sample. Single spot (or depth profile) ablation modes are recommended at a spot size of ~50 to 100 µm and a repetition rate of 10 Hz. Program the laser parameters.
Procedure | 9.3.4 | Requirement | Initiate the acquisition of the analytical signals using the ICP-MS software. Each data acquisition shall be comprised of a transient signal of intensity versus time for each element; each transient shall include 20–30 seconds of background (gas blank) measurement, followed by 50–60 seconds of ablation of the sample, followed by 10–30 seconds of post-ablation blank measurement.
Procedure | 9.4 | Requirement | Conduct replicate ablations at different locations within the fragment(s). Locations shall be spaced sufficiently to avoid possible debris from other ablation halos. Collect replicate measurements to ensure that the questioned glass fragments and known glass source(s) are adequately characterized. Analyze a minimum of three replicates on each questioned sample examined and nine replicates on known glass sources. NOTE 1: It is recommended that the entire sequence be completed in a single session (a single day). As an example, when a known sample with 3 fragments is compared to a single recovered fragment, the analytical sequence may be as follows: (1) Calibration standard (1st, 2nd, and 3rd spot) (2) “Known” fragment #1 (1st, 2nd, and 3rd spot) (3) “Recovered” fragment (1st, 2nd, and 3rd spot) (4) Calibration verification standard (1st, 2nd, and 3rd spot) (5) “Known” fragment #2 (1st, 2nd, and 3rd spot) (6) “Known” fragment #3 (1st, 2nd, and 3rd spot) (7) Calibration standard (4th, 5th, and 6th spot) NOTE 2: A symmetrical arrangement of the analytical sequence of standards and samples is advantageous in minimizing the effects that may result from instrumental drift.
Procedure | 9.5 | Requirement | Once the acquisition is completed, use the software of choice to integrate the transient signals and determine the element concentrations of the samples. Some data reduction software allows real time data reduction so the analytical data may be immediately checked after each ablation.
Procedure | 9.6 | Requirement | For each isotope, the transient signal shall be integrated, corrected for the gas blank and normalized to the 29Si internal standard signal. The normalized signals shall be converted to element concentration units using the calibration. The data reduction software may be selected by each laboratory. Because of particle size effects that may lead to undesired fractionation, it is recommended that the first 15 seconds or more of data collected during ablation not be integrated. See Fig. 1 for illustration of transient signal integration recommendation.
Limits Of Detection (LOD) And Limit Of Quantitation (LOQ) | 10 | Section Title
Limits Of Detection (LOD) And Limit Of Quantitation (LOQ) | 10.1 | Informational Text | Limits of detection and limits of quantitation were determined by laboratory analysis procedures prior to publication of this test method. Limits of detection (LOD) were determined for each element by measuring procedure blanks on two non-consecutive days using a quadrupole ICP-MS and a Nd-YAG 213-nm laser. Each day, the mean and standard deviation of replicate measurements were calculated. LODs were computed as 3 times the standard deviation. LOQs were computed as 10 times the standard deviation. Normalized intensities were converted to concentrations using the calibration data for each element for that day. Values reported in Appendix X1 are two-day averages.
Limits Of Detection (LOD) And Limit Of Quantitation (LOQ) | 10.2 | Informational Text | Table X1.1 gives the LODs, LOQs, and concentrations measured for sets of previously analyzed soda-lime glass samples. All values are expressed as µgg-1.
Calculation and Interpretation of Results | 11 | Section Title
Calculation and Interpretation of Results | 11.1 | Requirement | The procedure below shall be followed to conduct a forensic glass comparison when using the recommended match criteria:
Calculation and Interpretation of Results | 11.1.1 | Requirement | For the Known source fragments, using a minimum of 9 measurements (from at least 3 fragments, if possible), calculate the mean for each element.
Calculation and Interpretation of Results | 11.1.2 | Requirement | Calculate the standard deviation for each element. This is the Measured SD.
Calculation and Interpretation of Results | 11.1.3 | Requirement | Calculate a value equal to at least 3% of the mean for each element. This is the Minimum SD.
Calculation and Interpretation of Results | 11.1.4 | Requirement | Calculate a match interval for each element with a lower limit equal to the mean minus 4 times the SD (Measured or Minimum, whichever is greater) and an upper limit equal to the mean plus 4 times the SD (Measured or Minimum, whichever is greater).
Calculation and Interpretation of Results | 11.1.5 | Requirement | For each Recovered fragment, using as many measurements as practical, calculate the mean concentration for each element.
Calculation and Interpretation of Results | 11.1.6 | Requirement | For each element, compare the mean concentration in the Recovered fragment to the match interval for the corresponding element from the Known fragments.
Calculation and Interpretation of Results | 11.1.7 | Informational Text | If the mean concentration of one (or more) element(s) in the Recovered fragment falls outside the match interval for the corresponding element in the Known fragments, the element(s) does not “match” and the glass samples are considered distinguishable.
Precision and Bias | 12 | Section Title
Precision and Bias | 12.1 | Informational Text | An interlaboratory study was conducted in 2009 by the Elemental Analysis Working Group (EAWG). All analyses were conducted by LA-ICP-MS using either NIST 612 or the FGS calibration standards. Each of seven laboratories tested three standard reference glasses using 7 replicate sample measurements of NIST 1831, FGS1 and FGS2. The standard NIST 1831 is a soda-lime sheet glass manufactured by NIST. The standards FGS1 and FGS2 are soda-lime glasses manufactured by SCHOTT AG and distributed by the Bundeskriminalamt.
Precision and Bias | 12.2 | Informational Text | The bias and precision results for each of the glasses are tabulated in Appendix X2. The terms repeatability and reproducibility are used as specified in Practice E177. As suggested by this practice, the precision indices (95 % limits) were calculated by multiplying the respective standard deviation percentages by a factor of 2.8.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed |  | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics