## Instructions for Use
AAFS Standards Resources and Training
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions, an on-demand training webinar and instructional videos.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASTM E3296-22
ANSI/ASTM E3296-22
Standard Guide for Using Pyrolysis Gas Chromatography and Pyrolysis Gas Chromatography-Mass Spectrometry in Forensic Polymer Examinations
Discipline: | Trace Evidence
Checklist version: | V1.2
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Sample Handling | 5 | Section Title
General Considerations | 5.1 | Section Title
General Considerations | 5.1.1 | Informational Text | Practice E1492 as well as the relevant portions of the SWGMAT Trace Evidence Quality Assurance Guidelines and Trace Evidence Recovery Guidelines are followed for the general collection, handling, documentation, and tracking of specimens and samples.
General Considerations | 5.1.2 | Informational Text | Since PGC is a destructive technique, considerations are made prior to analysis to ensure that the entire sample is not consumed.
General Considerations | 5.1.3 | Recommendation | Prior to sampling, any foreign debris or contaminants that are visually observed by stereomicroscopy should be scraped away or otherwise removed.
General Considerations | 5.1.4 | Informational Text | Except where sample size or condition makes it prohibitive, samples for PGC or PGCMS analysis are prepared in a manner that permits resolution and analysis of individual layers. This allows the organic components from each layer to be effectively characterized and attributed to the appropriate layer. When samples are too small to isolate individual layers, samples from multiple layers can be analyzed simultaneously. Samples for comparison are prepared under the same conditions and in the same manner whenever possible.
General Considerations | 5.1.5 | Recommendation | Sample size can vary in accordance with polymer type, pyrolysis method, column type, chromatographic conditions, and detection methods. As a general approach, varying amounts of a known polymeric material can be analyzed to determine the minimum amount that provides sufficient signal to identify the characteristic components. Once optimal sample size is determined, samples of similar size and shape should be used for analyzing replicates and for comparison purposes. Instrument manufacturer recommendations and published studies of like materials can be used for guidance in determining approximate sample size.
General Considerations | 5.1.6 | Informational Text | Samples are positioned on or within appropriate sample holders to maximize pyrolysis. More specific guidance on sample holders is discussed in 6.7.1.
Fibers | 5.2 | Section Title
Fibers | 5.2.1 | Recommendation | Pyrolysis should be conducted on individual fibers, when possible. Optimal results are obtained from fiber lengths of 1–2 cm, although fibers as small as a few mm in length can be analyzed. As little as 25–30 µg has been reported to be sufficient for some acrylic fibers.
Paint | 5.3 | Section Title
Paint | 5.3.1 | Recommendation | Suitable paint sample sizes for pyrolysis range from 5–150 µg. Generally, more sample should be used when the layer is heavily filled with inorganic material.
Tape | 5.4 | Section Title
Tape | 5.4.1 | Informational Text | Suitable tape sample sizes for pyrolysis range from 10–60 µg, depending on the type of tape as well as the chemical composition of the material (for example, amount of inorganic filler, type of elastomer). A 0.5 mm square of electrical tape has been reported to correspond to approximately 25 µg.
Tape | 5.4.2 | Recommendation | Adhesive can be removed from tape for analysis using several methods. When preparing the adhesive for analysis, care should be taken to make sure any reinforcement materials (that is, fibers) are not present in the sample.
Tape | 5.4.3 | Recommendation | The tape backing can be isolated for analysis by removing the adhesive using a metal probe, a scalpel, an appropriate solvent, or a combination of these methods. Alternately, a thin peel of the tape backing can be used. The thin peel should only be used if the backing is known to consist of a single layer polymer.
Other Polymeric Materials | 5.5 | Section Title
Other Polymeric Materials | 5.5.1 | Informational Text | Suitable sample sizes for pyrolysis of other polymeric materials generally range from 10–100 µg, with automotive lens and rubber samples closer to 100µg and cosmetic samples approximately 25–50 µg.
Instrumentation and Operating Conditions | 6 | Section Title
Instrumentation and Operating Conditions | 6.1 | Informational Text | PGC instrumentation consists of two distinct components: (1) the pyrolysis unit, where sample fragmentation occurs, and (2) the gas chromatograph where separation and detection of the pyrolyzate fragments occur. In instances where a mass spectrometer is the detector, it is considered to be a third distinct part of the instrument, as it aids in the identification of selected pyrolyzates. Some systems include an interface between the pyrolysis unit and the gas chromatograph.
Instrumentation and Operating Conditions | 6.2 | Informational Text | The two most common detectors for use with PGC include the flame ionization detector (FID) and mass spectrometer (MS).
Instrumentation and Operating Conditions | 6.2.1 | Informational Text | The FID is capable of detecting a broad range of combustible pyrolyzates. It is relatively inexpensive and has a large linear dynamic range. Identifications and comparisons with data from an FID are based on retention times of peaks and the appearance of overall pyrogram patterns.
Instrumentation and Operating Conditions | 6.2.2 | Informational Text | The MS provides information about the individual fragments of the pyrolyzates, which enhances the ability to chemically classify the different polymer components. By selecting ions of interest, classes of compounds are able to be selectively viewed and searched for specific compounds.
Instrumentation and Operating Conditions | 6.2.2.1 | Informational Text | Various types of mass spectrometers are available for use as GC detectors.
Instrumentation and Operating Conditions | 6.2.2.2 | Recommendation | An electron ionization source is recommended instead of a chemical ionization source for applications discussed in this guide.
Instrumentation and Operating Conditions | 6.2.3 | Informational Text | Reconstructed total ion chromatograms in PGC/MS look similar to PGC chromatograms and provide comparable information to conventional PGC analysis.
Pyrolysis Systems | 6.3 | Section Title
Pyrolysis Systems | 6.3.1 | Informational Text | Three different types of pyrolyzers are available: (1) resistively-heated, (2) Curie point/inductively-heated, and (3) microfurnace.
Pyrolysis Systems | 6.3.2 | Informational Text | In resistively-heated pyrolysis, the coil heats first, then transfers the heat through the sample holder (for example, quartz tube) and heats the sample. Temperature ramping is possible and the desired temperature can be selected by controlling the current. The coil reaches its final temperature in 1–2 seconds; however, it can take several seconds for the sample to pyrolyze for a total wait time close to 10 seconds. Some drawbacks are that coil probes can exhibit “cold spots” and the coil is slow to heat the thermal mass of the quartz tube when compared with other pyrolysis systems.
Pyrolysis Systems | 6.3.3 | Informational Text | With inductively-heated or Curie point pyrolysis, the sample is either coated on or placed in a ferromagnetic wire, ribbon, or boat. The type of wire, ribbon, or boat that is used determines the maximum temperature that can be reached. Different alloy compositions are available for a wide range of temperatures. Temperature ramping is not an option with this technique.
Pyrolysis Systems | 6.3.4 | Informational Text | Microfurnace pyrolyzers use a quartz or metal cup to introduce the sample into the heated analytical chamber (microfurnace). The microfurnace is typically held at a single temperature without the need for equilibration; this allows for greater reproducibility than observed with other pyrolysis systems. The thermal mass of the cup is negligible compared to the microfurnace, so the cup and sample reach the set temperature nearly instantaneously
Pyrolysis Temperature | 6.4 | Section Title
Pyrolysis Temperature | 6.4.1 | Recommendation | The pyrolysis unit should use a temperature that results in complete pyrolysis without causing excessive bond breakage or fragmentation. Excessive fragmentation leads to a commonality of fragmentation data in the resulting pyrogram and difficulties in classification, identification, or discrimination.
Pyrolysis Temperature | 6.4.2 | Informational Text | The same analysis conditions (for example, set point, heating rate, time) are used for samples being compared.
Instrument Operating Parameters | 6.5 | Section Title
Instrument Operating Parameters | 6.5.1 | Informational Text | Operating conditions (for example, time, column selection, temperature programming) are optimized for pyrolysis, chromatographic separation, and detection. As the analyst determines specific analytical needs, actual operating conditions can vary. Variation of the operating values is permitted. The choice of conditions is based on the quality of pyrograms produced with regard to peak separation, resolution, and reproducibility.
Inlet Temperature | 6.5.2 | Informational Text | Typical inlet temperatures range from 200–300 °C.
Split Ratio | 6.5.3 | Informational Text | The sample introduced into the instrument is split so that neither the column nor detector are saturated by the sample. Typical split ratios range from 20:1 to 100:1, depending on sample size.
Instrument Operating Parameters | 6.5.4 | Informational Text | Oven parameters are critical for separation of pyrolyzates. Oven temperature, ramp rates, column type, and gas flow rates influence the pyrograms obtained. Molecular weight, resolution between peaks, potential for sample carryover, and total run time, are considered when selecting these parameters.
Instrument Operating Parameters | 6.5.4.1 | Recommendation | Temperature ramps should be used, as samples can contain a wide range of fragment sizes.
Instrument Operating Parameters | 6.5.4.2 | Informational Text | A defined temperature profile is used as well as a stable carrier gas flow rate.
Instrument Operating Parameters | 6.5.4.3 | Informational Text | Column carrier gas flow rates affect make-up gas flow rates. Therefore, for PGC with FID, manufacturers recommended flow rates are used for hydrogen and oxygen.
Instrument Operating Parameters | 6.5.5 | Informational Text | The choice of GC column is dependent upon the polarities of pyrolyzates. The use of two independent chromatographic columns (one high polarity and one low polarity) concurrently in a single PGC system can provide improved sensitivity and complementary pyrograms.
Example Instrumental Parameters | 6.6 | Informational Text | The following suggested operating conditions are meant as general guidance for PGC/MS analysis using a resistively-heated probe; however, parameters are optimized for each instrument. Additional examples of instrumental conditions are available in a variety of references. Pyrolysis temperature and time: 700 °C for 10 sec GC oven temperature program: Interface temperature: 275 °C Column: non-polar capillary column (30 m 0.25 mm ID) Carrier gas: Helium Pressure: 200 kPa Split flow ratio: 75:1 Oven program: Column remains at 40 °C for 2 minutes Ramp temperature: 6 °C/min to 295 °C Hold at 290 °C for 5 min Mass spectrometer: Scan speed: 1000 m/z per sec Time interval: 0.5 seconds Mass range: 50–500 m/z Transfer line: 290 °C Total run time: ~ 47 minutes
Quality Control | 6.7 | Section Title
Sample Holders | 6.7.1 | Section Title
Sample Holders | 6.7.1.1 | Informational Text | Reusable sample holders (for example, quartz tube, cups, or metallic ribbon) are cleaned before each use and have been demonstrated to be free of contamination on subsequent runs. One method for cleaning sample holders includes heating them in an apparatus (for example, hand-held propane torch, muffle furnace, filament coil) prior to use.
Sample Holders | 6.7.1.2 | Informational Text | Containers are discarded if they are damaged or significant residues are observed. Generally, the residues left behind are inorganic components that can be characterized using other techniques, if desired.
Sample Holders | 6.7.1.3 | Informational Text | For coil pyrolysis probes the coils are evenly spaced by gently adjusting them with forceps. All samples are placed in the same position inside the sample holder to ensure reproducibility. In the case of quartz tubes, samples can be held in position in the tube by the addition of quartz wool or a filler post.
Blanks | 6.7.2 | Section Title
Blanks | 6.7.2.1 | Informational Text | A system blank is run using all aspects of the system prior to every case sample to ensure that there is no contamination or carryover.
Blanks | 6.7.2.2 | Informational Text | The system blank is evaluated against acceptance criteria for blank runs in laboratory procedures. An ideal blank exhibits no peaks. Any peaks observed in a system blank are evaluated to determine the acceptability of the blank.
Performance Check | 6.7.3 | Section Title
Performance Check | 6.7.3.1 | Informational Text | For PGC/MS, the mass spectrometer is tuned using a standard reference compound (for example, perfluorotributylamine) to ensure optimization of peak shape, accurate mass assignments, and sensitivity.
Performance Check | 6.7.3.2 | Informational Text | Prior to casework samples, a standard polymer (for example, polyethylene, polystyrene, Kraton D1107) is analyzed using the same analytical conditions. The pyrogram produced by the standard is evaluated to ensure it meets the laboratory’s established quality control criteria such as demonstration of reproducible separation of peaks, peak area ratios, and pyrolysis fragmentation.
Identification, Comparison, and Interpretation | 7 | Section Title
Identification | 7.1 | Section Title
Identification | 7.1.1 | Informational Text | Identification of polymers by PGC or PGC/MS is based on the interpretation of pyrograms and comparison to polymer references. In combination, retention times, chromatographic peak patterns, and mass spectra can be used to identify and to compare pyrolysis products.
Identification | 7.1.2 | Informational Text | A sample pyrogram is compared to a reference library or a contemporaneously analyzed reference material to identify a polymer.
Identification | 7.1.3 | Recommendation | Optimal library pyrograms originate from the same instrument and protocol used in the current analysis. Traceable reference standards should be used in creating the library.
Identification | 7.1.4 | Informational Text | When MS is employed, individual chromatographic peaks can be identified by means of mass spectral library searches. The components identified could aid in determining the original starting materials of the manufacturing process.
PGC or PGC/MS Pyrogram Comparison | 7.2 | Section Title
PGC or PGC/MS Pyrogram Comparison | 7.2.1 | Recommendation | PGC or PGC/MS pyrogram comparisons should be conducted between pyrograms collected using similar sample preparations, similar sample characteristics (for example, thickness, topography), and similar instrumental parameters, as appropriate.
PGC or PGC/MS Pyrogram Comparison | 7.2.2 | Informational Text | Pyrograms are compared and interpretated based on the observation of any differences, or lack thereof, between the sets of PGC or PGC/MS data.
PGC or PGC/MS Pyrogram Comparison | 7.2.2.1 | Informational Text | Pyrogram overlay can be used for comparing data where the presence or absence of peaks, peak shapes, and relative intensities are all considered in the evaluation as to whether exclusionary differences exist between compared samples.
PGC or PGC/MS Pyrogram Comparison | 7.2.2.2 | Requirement | When assessing differences between pyrograms, consider sample limitations (for example, small samples, thin layers, dirty samples, sample smears that eliminate layer structure) and instrumental limitations (for example, sampling size, limits of detection).
PGC or PGC/MS Pyrogram Comparison | 7.2.3 | Informational Text | Possible reasons for pyrogram differences include dissimilar sample characteristics (for example, size, thickness, surface topography), heterogeneity, contribution from extraneous materials, or origination from different source materials. Additional samples can provide supplemental data to assist in assessing such differences.
PGC or PGC/MS Pyrogram Comparison | 7.2.4 | Informational Text | If suitable PGC or PGC/MS pyrograms are produced, comparisons can provide information regarding the potential relationship between the sources of the samples.
PGC or PGC/MS Pyrogram Comparison | 7.2.4.1 | Informational Text | When exclusionary differences are observed between compared pyrogram features, the sources of the samples are considered distinguishable by PGC or PGC/MS. Exclusionary differences in pyrogram comparisons: (1) are outside the variability of pyrograms originating from the same source; and (2) cannot be explained by considerations such as sample heterogeneity, contamination, different sample conditions, or different sample histories.
PGC or PGC/MS Pyrogram Comparison | 7.2.4.2 | Informational Text | When no exclusionary differences are observed between compared pyrogram features, the sources of the samples are considered indistinguishable by PGC or PGC/MS. Differences that are not considered exclusionary: (1) are within the variability of pyrograms originating from the same source; or (2) can be explained by considerations such as sample heterogeneity, contamination, different sample conditions, or different sample histories. If no exclusionary differences are observed in a PGC or PGC/MS comparison, samples can be analyzed by other analytical techniques to provide additional information about the potential relationship between the sources of the samples.
PGC or PGC/MS Pyrogram Comparison | 7.2.5 | Informational Text | PGC or PGC/MS pyrogram comparison is one part of a multi-analytical comparative approach. PGC or PGC/MS data alone can be used to distinguish the sources of compared samples but are otherwise not used independent of data obtained from other analytical techniques to reach an overall opinion regarding the potential relationship between the sources of the samples. An overall opinion that sources are indistinguishable is only reported when no exclusionary differences are observed in the analytical techniques that were applied.
Documentation | 8 | Section Title
Documentation | 8.1 | Requirement | Record the details necessary to support the interpretations made from each comparison.
Documentation | 8.2 | Recommendation | Case notes should be sufficient to allow an independent analyst to understand and evaluate all the work performed, independently analyze and interpret the data, and draw conclusions.
Documentation | 8.3 | Informational Text | A description of the evidence analyzed by PGC or PGC/MS, sample preparations, the order in which samples were run, the analytical instrumentation used, and its operating parameters are included in the case notes, case record, or otherwise recorded in accordance with laboratory procedures.
Documentation | 8.4 | Informational Text | Refer to Practice E1492, Guide E3260, and ISO/IEC 17025 for further guidance.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed |  | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics