## Instructions for Use
AAFS Standards Resources and Training
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions, an on-demand training webinar and instructional videos.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASTM E2808-21a
ANSI/ASTM E2808-21a
Standard Guide for Microspectrophotometry in Forensic Paint Analysis
Discipline: | Trace Evidence
Checklist version: | V1
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Sample Preparation | 6 | Section Title
Sample Preparation | 6.1 | Recommendation | The general collection, handling, and tracking of samples should meet or exceed the requirements of Practice E1492.
Sample Preparation | 6.2 | Requirement | Verify that the work area and tools used for the preparation of samples are free of all extraneous materials that could transfer to the sample.
Transmittance Measurements | 6.3 | Section Title
Transmittance Measurements | 6.3.1 | Requirement | Prepare hand-cut or microtome-cut thin sections for transmittance measurements. Although embedding followed by microtomy could be considered labor-intensive, this method (when questioned and known samples are mounted together) produces samples with an equivalent path length. Samples with equivalent path lengths are preferred for all transmission measurements and required for analysis of UV absorbers when the comparison relies upon differences in relative concentrations.
Transmittance Measurements | 6.3.2 | Informational Text | Paint sections of approximately 3 μm thick are generally appropriate for measurements of pigmented layers. Lightly colored samples or clear coats could require thicker sections (for example, up to 20 μm thick) for increased sensitivity of the instrument to low concentrations of UV absorbers or pigments.
Transmittance Measurements | 6.3.3 | Informational Text | Photometric reproducibility could be difficult to achieve with thin peels of paint layers or hand-cut cross sections due to uncontrolled path lengths within and between samples. This can lead to difficulties in the interpretation of shade (that is, intensity) differences when spectral curves do not exhibit exclusionary features.
Reflectance Measurements | 6.4 | Section Title
Reflectance Measurements | 6.4.1 | Requirement | Conduct specular reflection measurements on clean and undamaged sample areas. Clean the surface of the sample carefully with deionized water, or alcohol if necessary. Reflectance measurements are usually conducted without further preparation of the sample surface so that surface features (for example, weathering) that might distinguish between samples are not altered. In some cases, it will be necessary to remove surface features. Surface weathering, for example, can be removed by polishing the paint using a diamond paste or another polishing medium so the measured properties of the surface are comparable to data for new, undamaged paint surfaces. However, caution is recommended before removal of features as the potential for other evidential loss increases. Retain a portion of the original unaltered sample whenever possible.
Reflectance Measurements | 6.4.2 | Requirement | To obtain optimal results during spectral measurements, prepare the surface so it is free of scratches, dirt, and blemishes and mounted perpendicular to the optical axis of the microscope. Obtain a sharp focus on the paint sample surface.
Reflectance Measurements | 6.4.3 | Requirement | Conduct cross section reflectance measurements on polished or microtome cut surfaces. Prepare samples for comparison in the same manner. Simplify and enhance reproducibility of the surface finish by mounting and polishing edge-mounted known and questioned samples side by side.
Reflectance Measurements | 6.4.4 | Requirement | When the signal-to-noise ratio of a reflectance measurement is poor due to small sample size, prepare and analyze a thin cross section by transmittance.
Sample Preparation | 6.5 | Recommendation | For paint smears, different portions of the smear can be separated and mounted for analysis. These results can provide conclusive exclusionary data but could lack meaningful associative information. Smeared paints are difficult for reflectance measurements; analyzing smears by transmittance methods is recommended.
Mounting Samples | 6.6 | Section Title
Mounting Samples | 6.6.1 | Requirement | Mount and prepare questioned and known samples in the same manner.
Mounting Samples | 6.6.2 | Requirement | For reflectance measurements, record the angle of illumination and collection.
Mounting Samples | 6.6.3 | Requirement | For transmittance measurements, mount the thin section or smear particles on a microscope slide under a coverslip in an appropriate refractive index medium (typically one that is close to that of the sample).
Mounting Samples | 6.6.3.1 | Requirement | Use a microscope slide with transmittance characteristics appropriate for the region of the spectrum being analyzed. Glass (borosilicate) slides and coverslips are suitable for measurements in the visible and NIR portions of the spectrum. Mount samples to be analyzed in the UV region of the spectrum on quartz or fused silica; however, not all slides have the same UV transmittance characteristics.
Mounting Samples | 6.6.3.2 | Requirement | Add a mounting medium between the slide and coverslip for transmittance measurements. Select a mounting medium that is compatible with the sample (that is, will not dissolve the sample) and with the spectral range being investigated (that is, glycerol is transparent to the UV region of the spectrum). Mounting media include, but are not limited to, water, xylene, xylene substitutes, glycerol, refractive index oils (n = 1.52 or 1.66 are common).
Performance Checks | 7 | Section Title
Performance Checks | 7.1 | Requirement | Prior to use of the instrument, turn on the microscope, illumination sources, and spectrometer and allow them to stabilize. Allow the lamp to warm up and stabilize according to manufacturers’ instructions or laboratory experience, whichever yields consistent results.
Performance Checks | 7.2 | Informational Text | Checking instrument performance verifies that an instrument is operating within required standards and any errors that might affect data or analytical conclusions are known, limited, accounted for, or corrected. It is essential to demonstrate wavelength and absorbance/photometric accuracy through a performance check, such as that described in Practice E275.
Performance Checks | 7.2.1 | Requirement | Conduct a performance check daily, prior to analyses.
Performance Checks | 7.2.2 | Recommendation | A performance check should be conducted prior to use after any maintenance.
Performance Checks | 7.2.3 | Requirement | Use a similar configuration each time a performance check is conducted on the system to ensure that historical performance check data are comparable.
Performance Checks | 7.3 | Requirement | Maintain a record of all performance checks. A historical record of this data provides a mechanism for monitoring system performance and provides an operator with an early warning of system trends and deterioration.
Performance Checks | 7.4 | Informational Text | Performance check parameters include:
Performance Checks | 7.4.1 | Requirement | Wavelength Accuracy Routinely check wavelength accuracy over the measured range with the aid of National Institute of Standards and Technology (NIST)-traceable holmium, erbium, or didymium oxide filters. The resolution used during the wavelength accuracy checks should be the same or higher than that used in casework and consistent for each wavelength accuracy check. Transmittance is used for these measurements.
Performance Checks | 7.4.2 | Informational Text | Photometric Accuracy Neutral density filters are used to demonstrate the photometric response of the system is linear. A typical set of neutral density calibration filters could include some or all of the following filters: 0.1, 0.5, 1.0, 2.0, 2.5, and 3.0 absorbance units.
Instrument And Scanning Parameters | 8 | Section Title
Instrument And Scanning Parameters | 8.1 | Informational Text | Microscope parameters include:
Instrument And Scanning Parameters | 8.1.1 | Requirement | Illuminator Select an illuminator appropriate to the analysis being conducted. Select the emission spectrum of the illuminator that has sufficient intensity across the entire wavelength range of interest so as to provide a spectrum with an acceptable signal-to-noise ratio. Tungsten, halogen, and xenon are commonly used for visible and NIR analysis. Commonly, xenon lamps are used for UV analysis and mercury lamps are used for fluorescence excitation. While LED illuminators are available over much of the spectrum, they are of little utility for MSP due to their lower intensity and limited spectral range.
Instrument And Scanning Parameters | 8.1.1.1 | Informational Text | Instrument design with regard to lighting can affect the ease of collecting reflectance measurements. Epi-illumination results in specular reflectance of the illuminating light, causing a loss of detail in the viewed image, making layer distinction difficult to impossible at times. The illumination/measuring geometry in reflectance analysis of 45°/0° is preferred such that the specularly reflected portion is not gathered by the microscope. This geometry can be best achieved with the help of a darkfield/brightfield illuminator fitted with darkfield objectives. However, such illumination excludes UV measurements as UV transmitting darkfield objectives are not available. A general discussion of geometric considerations for reflectance and transmittance measurements is found in Guide E179.
Instrument And Scanning Parameters | 8.1.1.2 | Informational Text | Background/system/reference transmittance spectra can be used to monitor illuminator performance and warn of unsuitable system alignment.
Instrument And Scanning Parameters | 8.1.1.3 | Informational Text | Illumination Centration Slight adjustments to the position of the bulb can serve to increase or decrease the emission over specific regions of the spectrum. For example, it is possible to maximize UV illumination, often at the expense of some light in the visible wavelengths. Generally, the slight loss of intensity in the visible region is not problematic due to the high intensity of modern bulbs.
Instrument And Scanning Parameters | 8.1.1.4 | Recommendation | Illumination Intensity For some illuminators, this can be a fixed parameter. When the voltage of an illuminator is adjustable, it should be held fixed following the photometric intensity performance check (7.4.2).
Instrument And Scanning Parameters | 8.1.2 | Requirement | Field Diaphragm To minimize stray light, with the specimen in focus, the edges of the field diaphragm should be brought into view and sharply focused with adjustment of the substage condenser. The field diaphragm is then opened so that its edges are either just outside the collection aperture or just beyond the field of view. Readjust the focus and size of the aperture when the objective is changed.
Instrument And Scanning Parameters | 8.1.3 | Recommendation | Substage Aperture (that is, Condenser Iris) The substage aperture should be opened until the desired image contrast is obtained. As adjustment of this aperture impacts the amount of light reaching the detector, this aperture setting could be different from that used to produce an ideal image. In some instances, it is desirable to further increase the opening of the substage aperture to allow more signal to reach the detector. Aperture adjustment should occur prior to the collection of the background spectrum, and should remain in a fixed position between the background and sample spectra collection and between samples when they are being compared. The appropriate aperture level is typically that which produces an emission intensity for the most intense emission peak of no more than approximately 80 % of the detector saturation value. It is critical that the detector is not saturated anywhere over the region being measured.
Instrument And Scanning Parameters | 8.1.4 | Requirement | Objective Select an objective that permits imaging of the paint layer or inclusion to be analyzed. A balance between the objective magnification and size of the measuring aperture is selected by the analyst. Typically, measurements are made using objectives between 10× and 50×. Quartz optics are required for measurements made in the UV region. Once the appropriate objective is selected, measure all samples being compared at a fixed magnification.
Instrument And Scanning Parameters | 8.1.5 | Requirement | Measuring Aperture If the MSP system is equipped with variable or multiple collection apertures, select and use the same aperture for the measurement of all samples being compared. In general, select the largest aperture that will remain within the boundaries of the sample area to be measured. This is particularly important for paints with effect pigments, which show variation over larger areas. The edge of a layer or sample should be avoided due to edge effects that could impact the spectrum. For smears and thin peels, an oversized aperture (one that extends beyond the boundary of the sample) is undesirable in transmittance measurements as it increases the noise in the spectrum. In fluorescence emission, the spectrum is not the result of a ratio to a reference scan; the strength of the signal is determined by absolute counts. Therefore, an oversized aperture when analyzing smears or thin peels for fluorescence emission can be used to increase the signal reaching the detector. As the background is black, there is no significant increase in noise to detract from the quality of the collected data.
Instrument And Scanning Parameters | 8.1.6 | Requirement | Filter Cube Some systems have a filter cube turret. Filter cubes can include a blank cube for transmission measurements, a mirror for reflection, and combinations of excitation and barrier filters for fluorescence measurements. Use a filter cube appropriate to the measurement being made. Measure all samples being compared under a given configuration with the same filter cube in place.
Instrument And Scanning Parameters | 8.1.7 | Recommendation | Phototube Diverter Some systems have an adjustable phototube that permits light to be diverted to the ocular, spectrometer/camera, or split. Spectral artifacts (for example, interference fringes on the baseline) could be visible if the diverter is in the split position. The diverter should be set such that all light is directed to the spectrometer during measurements to ensure the highest possible signal to noise ratio.
Instrument And Scanning Parameters | 8.2 | Informational Text | Spectrometer parameters include:
Instrument And Scanning Parameters | 8.2.1 | Informational Text | Detector While the detector is not an adjustable variable post-purchase, it is a critical component during data collection. Two main detector types are available for use with MSP systems:
Instrument And Scanning Parameters | 8.2.1.1 | Informational Text | Photomultiplier tube (PMT) detectors consist of a photocathode, held at a positive potential, and a series of dynodes with successively lower potential to amplify the signal and convert photons of light into electrical energy. PMTs are sensitive, provide a high signal-to-noise ratio, and have good spectral resolution. PMTs are typically single channel detectors and are generally used in scanning spectrophotometers in conjunction with a grating, where the sample is scanned by stepping through wavelengths to create a spectrum point by point.
Instrument And Scanning Parameters | 8.2.1.2 | Informational Text | Semiconductor detectors are composed of a monochromator fitted with a diffraction grating and an array detector (for example, charged couple devices [CCD]) that acts as the photosensitive device. A CCD detector generally has a lower signal-to-noise ratio as compared to a PMT detector, but measurement time is drastically reduced because of simultaneous detection of the full spectral range. For this reason, CCD detectors are far more common than PMT detectors. The semiconductor detector’s resolution will depend on the number of pixels in the array, the dispersion and line spacing of the grating, and the distance between the grating and the array.
Instrument And Scanning Parameters | 8.2.2 | Requirement | Wavelength Range Select a wavelength range to reflect the spectral range to be measured and one that is appropriate to the instrument capabilities, sample, and sample preparation conditions as discussed in Section 6. This range typically falls between 190 and 1100 nm (transmittance UV-Vis-NIR).
Instrument And Scanning Parameters | 8.2.3 | Informational Text | Resolution The resolution is predominantly defined by the grating and slit size. For most commercial MSP instruments, the grating and slit size are fixed, which results in a fixed maximum resolution. For CCD detectors, pixel size and spacing also have an effect on resolution due to pixel binning. Pixel binning (that is, resolution factor) reduces the resolution and increases the signal to noise ratio. The resolution factor is selected as a balance between signal to noise ratio and spectral resolution.
Instrument And Scanning Parameters | 8.2.4 | Requirement | Integration Time Most MSP systems allow the user to define the integration time. Some software packages have a built-in functionality, which automatically adjusts the integration time. The integration time is determined from a background spectrum. When automatic integration time is not available, set the integration time such that the detector electronics are not saturated.
Instrument And Scanning Parameters | 8.2.5 | Requirement | Number of Scans Most MSP systems allow the user to define the number of scans to average for a single spectrum. Select the optimum number of scans that yields the desired signal to noise ratio. Generally, the signal to noise ratio is increased by the square root of the number of scans.
Sample Analysis | 9 | Section Title
Sample Analysis | 9.1 | Requirement | Prior to sample analysis, allow the instrument to stabilize and then conduct a performance check in accordance with Section 7.
Sample Analysis | 9.2 | Requirement | Optimize the illumination field, measuring aperture, and magnification for the sample under investigation. Experimental conditions and instrument settings (that is, objective, aperture size, lamp voltage, scan/spectrum averaging, spectral resolution, or sample focal plane) should be identical for known and questioned samples to minimize experimental errors and sampling variations.
Sample Analysis | 9.2.1 | Informational Text | Experimental conditions include the length of time that a sample resides in the illumination beam. Sample temperature can increase with examination time as the sample absorbs heat, which can cause irreversible degradation in some dyes and pigments (for example, bleaching). Minimizing the time the sample is illuminated with UV light helps reduce specimen heating and UV degradation during analysis.
Sample Analysis | 9.2.2 | Informational Text | Consistent sample conditions are especially critical when collecting fluorescence data because the amplitude of the resulting spectrum is directly correlated with the collection conditions (that is, there is no reference scan to ratio against).
Sample Analysis | 9.2.3 | Informational Text | Inherent to any grating based spectrophotometer, energy diffracted from grating rulings can be polarized (predominantly polarized perpendicular to the rulings) and can result in spectral artifacts. This effect (Wood’s anomaly) can be observed as a distinct transmittance maximum within the spectrum. Therefore, consistent or similar sample orientations are preferred when comparing samples.
Collection of Reference Spectra | 9.3 | Section Title
Dark Scans | 9.3.1 | Section Title
Dark Scans | 9.3.1.1 | Informational Text | Dark scan is a reference spectrum collected when the light from the microscope is blocked from the detector and is a measurement of instrument noise.
Dark Scans | 9.3.1.2 | Requirement | A new dark scan collection is required at least prior to and between microscope slide or sample preparations.
Dark Scans | 9.3.1.3 | Requirement | A dark scan is required for transmittance, reflectance, or fluorescence emission spectra.
Background/Reference Scans | 9.3.2 | Section Title
Background/Reference Scans | 9.3.2.1 | Informational Text | Background/system/reference scan is a spectrum that measures the light transmitting/absorbing effect of all the system components except the sample of interest (for example, light source, optics, microscope slide, cover slip, mounting media).
Background/Reference Scans | 9.3.2.2 | Requirement | The background scan is required for transmittance or reflectance measurements, but is not relevant to fluorescence emission measurements.
Background/Reference Scans | 9.3.2.3 | Requirement | For transmittance measurements, collect a new background reference scan for each new microscope slide and cover slip preparation, aperture size, and instrument configuration. To avoid heterogeneities in the mountant, collect a new background reference scan when moving outside the field of view on the same slide. In general, a reference scan collected before every sample scan will best compensate for the effects of adjusting the fine focus on the microscope.
Background/Reference Scans | 9.3.2.4 | Requirement | Use an appropriate white standard material (for example, BaSO4, MgO, a white unglazed ceramic tile, commercial grade opal glass, or sintered polytetraﬂuoroethylene [PTFE]) for the reflectance background reference scan. A new reference spectrum should be measured for each new reflectance sample as a precaution to avoid spectral errors caused by short-term instrument instability.
Sample Measurement | 9.4 | Section Title
Sample Measurement | 9.4.1 | Recommendation | At least five spectra should be collected from each non-effect paint film, each representing different sample areas or depths (that is, for clear coats). For paints containing effect pigments, a larger number of spectra will be required. The number of replicate analyses can be adjusted in an effort to capture the variation present within the sample. Small sample size or poor sample conditions can preclude the acquisition of sufficient spectra to meet these criteria.
Sample Measurement | 9.4.2 | Requirement | Attempt to select areas for analysis that are homogenous and flat. If possible, avoid large particles such as opaque effect pigments (for example, aluminum flake) or areas with obvious foreign debris.
Sample Measurement | 9.4.3 | Requirement | For reflectance analysis, select sample areas that are uniform in size and have minimal dust and scratches. If paint samples are measured in reflectance, and both samples are not oriented normal to the optical axis of the microscope, variations could be observed in their spectra. The reflectance differences are generally seen as variations in reflectance intensity rather than in curve slope(s) or absorbance cutoff point(s). Multiple sampling/scan setups could resolve these variations. Variations in spectra due to sample heterogeneity can be minimized by embedding samples side by side, then smoothing by cutting or polishing.
Sample Measurement | 9.4.4 | Requirement | Select known sample(s) from an area as near as possible to the suspected source point of the questioned sample(s). Differences could arise in measurements of paint samples from the same vehicle or other painted surfaces because of differences in sunlight exposure, weathering conditions, or repaired areas.
Spectral Comparison And Interpretation | 10 | Section Title
Spectral Comparison And Interpretation | 10.1 | Recommendation | MSP spectral comparisons should be conducted between spectra collected using similar sample preparation methods, similar sample characteristics (for example, thickness, topography), and similar instrumental parameters, as appropriate.
Spectral Comparison And Interpretation | 10.1.1 | Requirement | Compare the spectra and make interpretations based on the observation of any spectral differences, or lack thereof, between the sets of MSP data.
Spectral Comparison And Interpretation | 10.1.2 | Informational Text | Spectral overlay is a recognized method for comparing data where the presence or absence of peaks, peak shapes, and relative intensities are all considered in the evaluation as to whether exclusionary differences exist between compared samples.
Spectral Comparison And Interpretation | 10.1.2.1 | Informational Text | Comparison of known and questioned evidence can be conducted with compared spectra displayed in percent reflectance, percent transmittance, or absorbance formats, although certain information could be observed more readily in one format or the other. For example, absorbance data is directly proportional to concentration, and therefore, peak heights might be more critically scrutinized when comparing absorbance spectra from questioned and known samples with the same section thicknesses.
Spectral Comparison And Interpretation | 10.1.3 | Requirement | First and second derivative functions of the spectra can also assist in identifying inflection points and aid in the discrimination of samples. Effective use of derivative functions requires that spectra have low noise levels or high signal-to-noise ratios. Conduct spectral derivative calculations on absorbance data.
Spectral Comparison And Interpretation | 10.2 | Informational Text | Mean value spectra (that is, averaged) can be generated from replicate scans of each sample (for example, layer). Spectra are typically averaged and then compared to the averaged spectra of another item.
Spectral Comparison And Interpretation | 10.2.1 | Informational Text | Plots of standard deviation spectra (calculated from multiple spectra collected from a given sample/layer) can also provide a useful point of comparison. Standard deviation curves can be useful for estimating the known sample variation range; however, comparisons based upon standard deviation spectra use intensity as a criterion for comparison. Unless spectra are collected from sections prepared with the same thickness, the resulting standard deviation spectra have questionable value.
Spectral Comparison And Interpretation | 10.3 | Informational Text | The known and questioned sample comparison begins with the examination of the whole spectrum, followed by critical examination of each specific peak. The comparison includes examination of peak shape, minima, maxima, inflection points, troughs, shoulders, relative peak intensities, and the curves or slopes between peaks.
Spectral Comparison And Interpretation | 10.4 | Requirement | Some spectral differences can be subtle and visually difficult to discern. In these instances, chemometric analysis could help assess compared samples. To employ chemometric analysis, collected case data is processed (that is, pretreatment) and a series of mathematical and statistical methods (for example, Principal Component Analysis, Agglomerative Hierarchical Clustering, Discriminant Analysis) are applied. Chemometrics are best applied to large data sets, meaning a greater number of replicate analyses, large populations of relevant samples, or both are required. The need for large data sets could limit the value of chemometrics when comparing samples, as size and condition of submitted evidence could prevent a suitable number of replicate analyses for statistical evaluation. Regardless of the statistical methods used, validate the model prior to use on casework samples. It is noted that consensus has not been reached in the relevant scientific literature on the most appropriate data pretreatment(s) or statistics for application to MSP data.
Spectral Comparison And Interpretation | 10.4.1 | Requirement | When assessing differences between spectra, consider sample limitations (for example, small samples, thin layers, dirty samples, sample smears that eliminate layer structure) and instrumental limitations (for example, sampling size, limits of detection).
Spectral Comparison And Interpretation | 10.4.2 | Informational Text | Possible reasons for spectral differences include dissimilar sample characteristics (for example, size, thickness, surface topography), heterogeneity, contribution from extraneous materials, or origination from different source materials. Additional samples can provide supplemental data to assist in assessing such differences.
Spectral Comparison And Interpretation | 10.5 | Informational Text | If suitable MSP spectra are produced, comparisons can provide information regarding the potential relationship between the sources of the samples.
Distinguishable Sources | 10.5.1 | Informational Text | When exclusionary differences are observed between compared spectral features, the sources of the samples are considered distinguishable by MSP. Exclusionary differences in spectral comparisons: (1) are outside the variability of spectra originating from the same source; and (2) cannot be explained by considerations such as sample heterogeneity, contamination, different sample conditions, or different sample histories.
Indistinguishable Sources | 10.5.2 | Informational Text | When no exclusionary differences are observed between compared spectral features, the sources of the samples are considered indistinguishable by MSP. Differences that are not considered exclusionary: (1) are within the variability of spectra originating from the same source; or (2) can be explained by considerations such as sample heterogeneity, contamination, different sample conditions, or different sample histories. If no exclusionary differences are observed in an MSP comparison, samples can be analyzed by other analytical techniques to provide additional information about the potential relationship between the sources of the samples.
Spectral Comparison And Interpretation | 10.6 | Informational Text | MSP spectral comparison is one part of a multi-analytical comparative approach. MSP data alone can be used to distinguish the sources of compared samples, but otherwise is not used independent of data obtained from other analytical techniques to reach an overall opinion regarding the potential relationship between the sources of the samples. An overall opinion that sources are indistinguishable is only reported when no exclusionary differences are observed in the analytical techniques that were applied.
Documentation | 11 | Section Title
Documentation | 11.1 | Requirement | For all analytical data, include a unique sample designation, the operator’s name/initials, and the date of analysis on all paper documents and electronic files that are retained as part of the case file.
Documentation | 11.2 | Requirement | Include in the case file the analyst’s assessment of the MSP data and all instrumental data that was used to reach a conclusion. Provide spectra in color or in a format such that spectra from various samples plotted together can be attributed to a legend when viewed in grayscale.
Documentation | 11.3 | Requirement | Include in the case notes a description of the evidence analyzed by MSP and a description of sample preparation including any mounting medium used, the analytical instrumentation used, and the optimized operating parameters (that is, aperture size, objective, scan/spectrum averaging, spectral resolution).
Documentation | 11.4 | Requirement | When making comparisons, record in the case notes the analyst’s assessment of the MSP data.
Documentation | 11.5 | Requirement | When chemometrics are applied, record the data analysis method(s) and all parameters (for example, software name and version, confidence intervals) necessary to review the result.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed |  | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics